# What I wanted you to consider before the finalizing the design of the EAD Device

# I'm adding this after reviewing your EAD device system documentation sent to me earlier

I recently attended a seminar where some of the discussion around gait assessment gave me a few important clinical points I want to build into the device from the start, rather than realising I need them later.

The device is meant to buzz harder/vibrotactile feedback when a patient's foot moves the "wrong" way, to help retrain their walking after a stroke. For that to actually work safely, it needs to tell a real mistake apart from the sensor just being a bit noisy or off. The four things below are what I need to make that possible — to me, this is the heart of the whole device, not extras.

# What I want, based on your EAD system Documentation

| From your quotation | What I additionally want |
| --- | --- |
| Gait speed, cadence, step symmetry | I want these numbers to stay accurate for the whole session, not getting drift as the patient keeps walking. |
| 6-channel vibrotactile feedback for error augmentation | I want the vibrotactile feedback intensity to be driven by a per-step deviation score — a single quantitative value representing how far that step deviated from the target gait pattern — rather than a fixed-amplitude output or buzz |
| Real-time graphs (angles, symmetry, events) | I want the device to compare each patient to their own normal walk, not a generic one. |
| Data export (CSV, MATLAB format) | I want the full raw sensor readings exportable, with a timestamp on each one, not just the summary numbers. |

# 1. I want the numbers to stop drifting- I want Fixed sensor drift (ZUPT)

I got to know that the foot sensor will slowly lose accuracy the longer someone walks — a bit like a watch that runs a little fast over time. I'd like a fix built in so that every time the foot is flat on the ground for a moment, the system resets itself. Without this, I won't trust the walking-speed and distance numbers by the end of a session.

# 2. I want one clear number that says how "off" each step was- ie  One clear “error” number per step

This is the part I care about most. I need one simple number per step that tells me how far off that step was from normal, so the buzz has something real to react to. I want the buzz to actually get stronger the more off the step was, with a limit so it never gets too strong to be unsafe.

# 3. I want the device to learn each patient's own normal walk- Calibrate to each patient

I don't want patients compared against a standard healthy-population gait pattern — every patient walks differently, especially after a stroke. I'd like a short calibration walk at the start of each session that sets that patient's own baseline and saves it for next time." So the device should be enabled to compare the patient to their own normal walk, not a generic one — otherwise it flags things that aren't real errors

# 4. I want raw sensor data export — full IMU output (accelerometer, gyroscope, orientation) at the native sampling rate, with a synchronized timestamp on every sample, not just the processed summary metrics."- Export raw data + a timestamp

I'll need this later to check the device's accuracy against another system, for my research. I'd like the software to let me export the full raw readings, not just the summary numbers, with a timestamp on each one. So we can later check the device's accuracy against a reference system/validating system.